Description:	If replacing the “old” style PC-AT motherboard in an LS140 with the “new” style requires a different PROM to be installed in the 724-F-5.

Note:  “Old” style motherboard has locations for 4 SIMM’s (memory modules) while the “new” style has locations for 8 SIMM’s.  “New” PROM will work for either version of motherboard.

Procedure:

- Change motherboard in accordance with Sect. 5.26 of LS140 Service Manual.

- Replace PROM at location J5021 on 724-F-5 with P/N 205-910-010 if unit is a standard (2K) unit or with P/N 205-913-010 if unit is an L1 (20K) unit.

- Make sure that new motherboard has same number of SIMM modules as board being replaced. Standard units have 2 SIMM modules, L1 units have 4 SIMM modules.