Description:	This document contains a fix for a video problem in the field.  There is a problem with the connector from the CRT driver board and the power interconnect board (724-F-7).  When the connector becomes resistive the power supply to the CRT driver board oscillates and causes the high voltage to oscillate.  When the high voltage oscillates the video picture size also oscillates (“breathes”/”flickers”).

Parts Required:

None

Procedure: for ECO 1001 on 724-F-7 board

- Remove the cabinet of the LS140.

- Remove the 724-F-7 board (power supply inyerconnect board).

- Remove the two pin connector at location J7011 on the 724-F-7 board.

- Cut the connector off of the Yellow and Black wires that connected to connector J7011.

- Solder the Yellow wire into the hole where Pin 1 of connector J7011 was.

- Solder the Black wire into the hole where Pin 2 of connector J7011 was.

- Re-install the 724-F-7 board and cabinet.