Description: This procedure will describe the steps necessary to install LS-NET into an LS140.

Warning:  Disassembly described in this section is performed with power removed from the instrument.  Do not remove or replace any of the circuit board assemblies without first verifying that power has been removed from the instrument.  Anti-static protection measures should be taken prior to the handling of any printed circuit boards.

Procedure:

1.	Disconnect power cable from the instrument.

2.	Place the LS140 face down into the plastic front panel protective cover that was shipped with the unit or on a soft surface to protect the BNC’s and front panel knob.

3.	Remove the four screws attaching the rear feet to the rear bezel.

4.	Remove the six screws attaching the cabinet to the rear bezel.

5.	Hold the cabinet by the two sides.  Carefully lift the cabinet using a left and right rocking motion.

6.	Remove the screw holding the blank panel into one of the open slots on the rear casting.

7.   There are no switches or jumpers on the NET board to set, however the board must be software configured at the factory before shipping.

8.	Install NET board into open slot.

9.	Secure with one screw to rear casting.

10. If the unit is a 2K model, install the RAM IC that was enclosed with the NET board upgrade kit on the 724-F-5 board in location J2025.  If the unit is a 20K model, (Option L1) proceed to next step.

11.	Reassemble cabinet.

12. Insert the appropriate network installation disk (LS140-2K NET or LS140-L1 NET) into the floppy drive and cycle power.  Wait for the Installation Complete message.  Remove disk form disk drive.

11.	Connect keyboard to LS140 through hole in rear panel of unit.

12.	Insert EE_EDIT disk into floppy drive and power up unit.

13.	Press F7 to add record.

14.	Type in what appears on barcode of new NET board exactly as it appears. (Include spaces).  Press enter.  System will ask for last two digits of ECO number. Type 00 enter.

15.	Remove disk from floppy drive and cycle power. Installation complete.

16. Inform T&M/NY service department of upgrade and unit serial number.