Purpose: There are two different versions of 724-F-5.  If the board is in a unit with 2K acquisition memory size then the firmware is stored in battery backed up RAM on the 724-F-5 board.  If the unit has 20K acquisition memory size then the firmware is located on the hard drive.  Also, the boot PROM is different for these two different types of units.

Parts Required: 	724-F-5

Boot PROM’s	(Shipped with 724-F-5)

Zero Ohm resistor / Buss Wire

Description:

- Determine the memory configuration of the unit being repaired.  A 20 K unit will have a hard drive mounted underneath the floppy drive.

- If the unit is an L1 (20K) then the PROM on the 724-F-5 board should be 205-913-010.  If it is a standard 2K then the PROM should be a 205-910-010.

- If the unit is an L1  then locations J5022, 5023 and 5024 should be empty.  If it is a standard unit then these locations should have RAM chips (205-298-128).  These can be removed from the board that is being replaced. Note:  IF the unit is a standard unit and has the network option installed, then there will need to be a RAM at location J5025. This can also be taken from the board being removed.

- If the unit is an L1 then there should be a zero ohm resistor (or buss wire) shorting location R5008.  If it is a standard unit then this location should be open.

- If the unit is a standard unit, the firmware will need to be installed after completing this installation.