This versatile instrument has a wide range of application.  It will measure voltage, current and resistance values accurately, and the design incorporates an efficient output meter.  Following is a description of how the various sections are used.  There are further and wider applications for this instrument which will manifest themselves as the operator becomes more familiar with the Multimeter.

Unless the operator understands the voltage and current readings of various circuits, it is advisable to always use the highest range available to obtain an approximate reading, and then choose a lower range which will be more suitable for an accurate reading.  This will prevent damage to the meter from excessive overload.

D.C. YOLTAGES

Turn the central selector switch to the desired voltage range and make sure that the right-hand switch is turned to that position marked "D.C."  The negative, or black, test lead is inserted in the negative jack on the instrument, and the red test lead is inserted in the positive jack.  The two test prods are then touched to the necessary parts of the apparatus under test, and the meter will read the difference in potential between the two points touched, which is actually the voltage.  It is necessary to remember that voltage is the difference in potential between any two points.

If it is desired to measure the voltage on the elements of a valve, the metal chassis of a radio receiver or amplifier is usually regarded as forming the negative side of the circuit, and the elements concerned as forming the positive side of the circuit.   For instance, if it is desired to measure the plate voltage of a valve, the appropriate range would be selected, the positive test lead placed on the plate contact of the valve, and the negative test lead placed on the chassis.  The meter would then read the valve's plate voltage.  This method does not apply to the measurement of negative grid bias.

To measure the negative grid bias, the negative test prod is placed on the grid and the positive test prod is placed on the negative filament or cathode contact. The negative bias will then be indicated on the meter.  This method will be inaccurate if a high value of resistance is included in the grid circuit, such as a resistance capacity coupled stage.  In this case, the negative test prod should be placed on the end of the grid leak resistor, which does not connect to the grid.

When making voltage measurements, it is not necessary to remove or disconnect any wires.

A.C. VOLTAGES

To measure alternating voltage, the only rearrangement of the controls on the instrument is to turn the right-hand switch to that position marked "A.C."  The appropriate voltage range is then selected in the ordinary way on the range selector switch, and the test prods, when plugged into the instrument, can then be connected to the two points between which it is desired to measure the voltage difference. Since alternating voltage has no fixed negative or positive potential, the negative or positive test lead from the instrument can be placed on either of the two points which are under test.  However, to form a safety habit, it is always wise to place the negative lead on the low potential side of the circuit or that side of the A.C. voltage which is connected to earth.  If this is inconvenient, the operator need not worry any further.

When measuring alternating voltages on the 10 volt range, the lowest meter scale, marked "10 V. A.C. only" should be used.  When using the 50, 250 and 1,000 V ranges, measurements should be made on the upper set of voltage graduations.

D.C. CURRENTS

In making current measurements. it is necessary to break the circuit and insert the test leads so that the meter is placed in series with the circuit.  For instance, to measure the plate current of a tube, the wire on the plate contact would be removed and connected to the positive side of the meter.  The negative meter lead would be connected to the plate contact and the selector switch would be turned to the desired range, and then the set switched on.  The plate current of the valve Would be registered on the meter.  This procedure also applies to any other circuit in which it is desired to measure current in milliamperes.  The circuit is simply broken and the meter inserted in the break to complete the circuit again.

Where the current value is unknown, it is always wise to commence on the highest range, and then turn the selector switch down to that range which gives the most convenient deflection of the needle on the meter.

For 10 amperes D.C. use the sockets marked "10 Amps" and turn control switch to 10 A.  When measuring D.C. currents make certain that the right hand switch is turned to position labelled "D.C."

A.C. CURRENTS

The MVA/2 by itself is only intended to measure alternating (A.C.) current on the 1 mA range, in which case the upper voltage graduations are used.  This range can be extended by using the MRCT "University" current transformer which is available as an extra.  The MRCT needs no soldered connections-it simply connects by terminals to the MVA/2 and a switch on the MRCT selects the desired A.C. current range.  It extends the MVA/2 to read the following A.C. current -2.5, 5, 10, 25, 50, 100, 250 and 500 mA A.C. and 1, 2.5, 5 and 10 amperes A.C.

RESISTANCE

This instrument will measure values of resistance in four convenient ranges 0-1,000 ohms, 0-10,000 ohms, 0-100,000 ohms, and 0-1 megohm.

To measure values of resistance below 1,000 ohms, the selector switch is turned to the position marked "R X 1"  The test leads are inserted in the instrument, and then the test prods are touched together so that the meter needle will swing right over to the position marked "0" on the upper meter scale.  If it does not exactly reach the "0" mark, the ohms compensator, at the left-hand side of the instrument, is turned until the needle indicates zero resistance.  The meter is then ready for use.

To measure resistance, one side, or both, of the resistance or other part should be disconnected from the rest of the circuit, and the test prods placed on its terminals.  The value of resistance will he shown on the ohms range.

For values up to 10,000 ohms, the switch is turned to the position marked "R X 10" and the scale figures must be multiplied by 10 to give the correct resistance.   For example, if you are measuring a resistance of 4,000 ohms, and the switches are turned to the correct position, then the meter needle will indicate 400.  Multiplying this by 10 gives 4,000, which is the correct reading, assuming that the resistor is in good order.

When measuring in the range of 10,000 ohms, it is necessary that the prods are touched together again and the needle adjusted for zero resistance by use of the ohms compensator.

For values up to 100,000 ohms, the range switch is turned to the position marked "R X 100," and the procedure is carried out as explained previously. For measurements up to 1 megohm, turn switch to "R X 1,000," and proceed as before.

In measuring resistance, it is necessary that the right-hand switch be turned position marked to the	"D.C."  Always, before measuring resistance, make certain that the test prods are touched together and the ohms compensator adjusted, so that the meter reads zero before operation.  The purpose of this ohms compensator is to compensate for any variation in battery voltage, which will enable you to obtain a maximum life from the built-in batteries

CAUTION-Before attempting to measure the resistance of any part of radio or electrical apparatus, be sure to switch off the power, or to disconnect one wire from each battery in the case of battery-operated equipment.

OUTPUT METER

In addition to measuring ordinary A.C. voltages over a wide range, the Multimeter can also be used as an output meter.  The right-hand knob on the instrument is turned to the position marked "OP," and the range selector is turned to an appropriate voltage and db range.  The test leads are inserted in the instrument, and one lead is attached to the chassis, while the other lead is touched to the plate of the output or power valve in the receiver or amplifier under test.

Small push-on clips are provided with the instrument.  These easily and conveniently fit on the test leads, so that it will not be necessary for the operator to hold these on to the point under check in the chassis.  They can be clipped on to any convenient wire or terminal, leaving the operator's hands free for alignment of the set.

If the range selector is turned to 10 volts when using this as an output meter, it will give a very sensitive reading  However, it will be found necessary for the volume control of the receiver to be kept low, so as not to damage the meter.  This 10-volt range is recommended for aligning sets.  If the output meter is required for a purpose other than alignment, the 50-volt or 250-volt will be found quite suitable.

Used in this manner, the instrument will facilitate the alignment of a receiver, especially when a modulated oscillator or signal generator is used as the source of signal.

The meter scale is also calibrated in db.  When using db ranges, the figure indicated by the Selector Switch must be added to that indicated by the meter pointer.

BATTERY REPLACEMENT

The resistance measurement section of this instrument utilises a standard 1.5-volt 950 dry battery cell in conjunction with three type 703 dry batteries. These usually last up to nine months without replacement.  It will be known when the battery is due for replacement by the fact that the ohms compensator on the panel will not enable the pointer of the meter to be brought right to the zero mark.

To replace batteries, remove the four screws on the edge of the instrument, and the batteries will be seen held in special clips inside the case.  Unsolder the leads from each end of the 703 batteries and replace with new batteries in exactly the same position.  Make sure that the lugs are soldered on to the new batteries in the same manner as they were to the old batteries.  The type 950 battery is held in a clip at the side of the meter, and can be quite easily removed without any further unscrewing and soldering.  When replacing cell, be sure that the brass cap is facing the front panel.

GENERAL

The primary purpose of this instrument is to measure D.C. voltage, currents and resistances, as well as A.C. voltages and output voltages.  The instrument is accurate, and is easily portable.

It will cover nearly all of the routine checking required in a radio receiver, and in general radio equipment.  It must be remembered that voltage measurements in a receiver will not only indicate that there is voltage available, but if they are measured through any of the components in the receiver, they will indicate whether that component is open circuit or otherwise by the indication of voltage on the meter.  Faulty and intermittent transformers can be tested by using the ohms section of the instrument, and continuity tests of all components can be made with this section.  The instrument will also indicate short circuits and open circuits, leaky condensers, faulty resistance, broken connections and incorrect values. The  output meter section can be used to indicate a standard of output from apparatus under test, and constant. use of the instrument will bring many new suggestions to mind.

The meter needle should normally coincide exactly with the zero end of the voltage and milliamp scales when the meter is not in use.  Should it not do so, it can be adjusted to zero by means of the large bakelite screw on the face of the meter, just below the glass.

“UNIVERSITY”

MULTIMETER AND

OUTPUT METER

MODEL MVA/2

INSTRUCTIONS FOR USE

UNIVERSITY GRAHAM INSTRUMENTS PTY. LTD.

For all measuring and Testing Instruments

5 NORTH YORK ST., SYDNEY